## Slide 1: ĐO CÁC THÔNG SỐ LINH KIỆN LED
Vũ Doãn Miên, Hà Nội 2013
Nội dung:
Đo các thông số quang
2. Đo các thông số điện
3. Đo các thông số nhiệt

## Slide 2
Các cấp độ linh kiện và thiết bị chiếu sáng LED

## Slide 3
Đo các thông số quang
Radiant Intensity – Cường độ chiếu sáng
Irradiance – Độ chói
Radiant power – Công suất quang
Radiance – Độ sáng
Luminous intensity – Cường độ sáng
Illuminance – Độ rọi
Luminous flux - Quang thông
Luminance – Độ chói
Tương quan giữa Radiometric và Photometric measurements
V() – spectral luminous efficiency for photopic response of the human eye

## Slide 4
Quang thông
Năng lượng bức xạ của một vật được xác định bằng công suất bức xạ và được đo bằng Watt. Phần năng lượng bức xạ trong dải phổ ánh sáng, tức là phần bức xạ trông thấy được bằng mắt thường được gọi là quang thông. Hay nói cách khác quang thông là đại lượng biểu thị phần công suất của bức xạ quang có tác dụng với mắt người qua một tiết diện nào đó. Đơn vị đo quang thông là lumen viết tắt là Lm, ta có 1W= 683Lm, ( đối với bước sóng λ=0,555 µm)
Cường độ sáng
Cường độ sáng của một nguồn sáng là quang thông của nguồn sáng đó bức xạ theo phương đã định, trong một đơn vị góc khối. Nó đặc trưng cho mật độ không gian của quang thông. Cường độ sáng = quang thông/ góc khối
Độ chói
Là đại lượng chỉ mức độ sáng của vật bức xạ ánh sáng, phản xạ ánh sáng, hoặc cho ánh sáng đi qua. Trong đo lường quang học độ chói chỉ được dùng với vật có kích thước hữu hạn, không dùng với nguồn sáng điểm. Độ chói của một nguồn sáng mặt được xác định bằng cường độ sáng theo hướng quan sát bức xạ từ một diện tích của nguồn sáng mặt mà hình chiếu của diện tích đó trên mặt phẳng vuông góc với hướng quan sát có trị số bằng 1 đơn vị diện tích.
Độ rọi
Đối với những vật thể không bức xạ ánh sáng mà được chiếu sáng, thì đánh giá mức độ sáng bằng độ rọi. Độ rọi trên một mặt thoáng nào đó được xác định bằng tỷ số quang thông rọi lên một đơn vị diện tích của mặt thoáng đó. Độ rọi trên một mặt thoáng tỷ lệ thuận với cường độ sáng của nguồn sáng và tỷ lệ nghịch với bình phương khoảng cách

## Slide 5
Radiant Flux: (Radiant Power)
Radiant flux φ also referred to as the power P, is the rate dθ/dt at which the radiation field
transfers radiant energy from one region to another. Where θ is defined as the radiant energy.
The unit of radiant power is the watt [W].
Solid Angle:
The solid angle Ω, in units of steradians [sr] subtended by a cone, is given by the expression
A/r2 where A is the area intercepted by the cone on the surface of a sphere of radius r
centered on the cone vertex.
Radiant Intensity:
The radiant intensity I is the flux per unit solid angle emitted by an entire source in a given
direction. It is expressed as I = dφ/dΩ in units of watts per steradian [W/sr]. A detector with
an active area A positioned at a distance r from a light source measures radiant power. This
configuration assumes a point source, therefore the inverse square law holds true. Distance r
and the detector area dA define the solid angle dΩ.
Irradiance:
Irradiance E is the flux per unit area received by a real or imaginary surface element and is
expressed as dφ/dA. The unit of irradiance is [W/m2]
Lumens measure light output at the source, while candelas measure the light emitted per unit of solid angle, a quantity that conveniently does not vary with the distance from the source of the light.
1 candela = 1 lumen/steradian
The surface area of a sphere is 4πr2,
The surface area of a steradian is just r2.
So a sphere measures 4π steradians, or about 12.57 steradians (1 sr  8% sphere)
Luminous intensity (cd) = Luminous flux (lm) / solid angle (sr)

## Slide 6
Các thông số quang cơ bản

## Slide 15
In case of lower precision photometers standard LEDs of same colors are needed as the test LEDs.
Substitution method would still allow high precision measurements.
Particularly important for blue and red LEDs

## Slide 26
Spectral responsibility of Si Photodiode (without filter)
Theoretical V(λ) function, relative spectral flux output of a blue LED, relative spectral distribution of a typical tungsten incandescent lamp (dash line), typical response of photopic detector (such as Si Photodiode with filter) (the different from V(λ) curves f’ 1 < 5%)

## Slide 27
Flux Measurement
• Integrating sphere size should be 20 cm diameter or more
• White interior coating reflectance should be 90% or higher (reducing errors due to different spatial power distributions)
• Photometer head shoud have a cosine response
Intensity measurement

## Slide 28
Intensity for point source does not depend on direction, size of aperture of detector
Polar plot of point source
Point source

## Slide 29
Polar plot of Lambertian source
Polar plot of some LED sources
Intensity for LED source depends on direction, size and shape of detector aperture
Lambertian source

## Slide 31
Intensity (I) angular distribution of various point light sources with rotational symmetry can be approximated as:
I = Io cosg-1θ
Where Io is the intensity normal to the source and g ≥ 1.
g = 1 for an isotropic point source
g = 2 for a Lambertian point source
g > 30 for an LED point source (LED lamp)
Illuminance E at a point on a surface (a calibrated detector):
E = I / d2
Whre d is distance between the source and the point (detector).
For a typical light bulb or luminaire, the far-field distance
"d" is defined to be at least ten times the maximum dimension of the source. For an LED, according to industry experts, no such rule-of-thumb exists and, the far-field distance "d“ can only be determined by finding the 1/r2 irradiance fall off region through measurement.
Spatial distributions of the radiation emitted by LEDs

## Slide 33
The drive current through an LED must be controlled. High current densities within the junction of the chip cause partial overheating which damages the crystalline structure of the LED die. At these areas are so called darkline defects, where light ceases to be generated.
This should not be confused with the maximum junction temperature of the LED, which is related to the higher temperature at which the optical grade epoxy starts to expand rapidly and increases the risk of catastrophic failures, such as broken wires or lifted LED dies.
The emitted light is a function of the forward voltage Vf and the forward current If
Electrical Properties
Operation at Constant Current
Modulated or Multiplexed Operation
One efficient method of driving LED devices is to pulse them (usually 100Hz to 1000Hz) with a high peak current for short durations on a low duty factor. This technique is frequently used to multiplex a number of individual LED lamps
Single-Shot Operation
During production, the measurements commonly made to characterize an LED are often carried out under single-shot operations

## Slide 34
Các phương pháp và thiết bị đo kiểm LED

## Slide 35
MEASURE:
Total Spectral Flux
Luminous Flux
Radiant Flux
Chromaticity
CCT and CRI
Peak Wavelength
Dominant Wavelength
I, V, and Luminous Efficacy
Thermal: Case Temperature Control vs. Electrical and Optical Parameters
APPLICATIONS:
Packaged LEDs
Modules & Arrays
Backlight Displays
Solid State Lighting
ILLUMIA®PRO LIGHT MEASUREMENT SYSTEMS (Labsphere com., USA)
CÁC THIẾT BỊ ĐO LED

## Slide 37
Sơ đồ khối máy đo LED
LED Tester SSP6612, Starspec com., China

## Slide 38
Các bộ phận chính của thiết bị

## Slide 39
Basic Test - Phép đo cơ bản

## Slide 40
Luminous test: - Đo kiểm quang thông theo dòng

## Slide 41
I-V test – Đo kiểm I-V

## Slide 42
Distribution of Light Intensity Angel Test – Đo kiểm phân bố cường độ ánh sáng theo góc

## Slide 43
Một số thông số của phép đo phân bố theo góc

## Slide 44
Database and Report Form - Cơ sở dữ liệu và dạng báo cáo
Cơ sở dữ liệu

## Slide 45
Dạng báo cáo kết quả đo